## Slide 1
COMPENSATION
FITC
FITC -
PE
PE -
FITC
-
-
+
+
gain=1
gain=1
detector
detector
PE
FITC
FL1
FL2
FL2-%FL1
Notes: Because fluorochromes overlap in their spectral emission it is not possible to just use optical filtration for their resolution. Electronic compensation provides a method for objective spectral resolution of fluorochromes. The principle is illustrated here. Cells or microspheres are labeled with a single color and evaluated. If FITC particles are used, both the FITC and the PE detectors see emission. If the FITC detector output is directed to the negative side of the PE op amp (bottom), it can be adjusted (FL2-%FL1) to exactly equal the height of the signal seen by the PE detector and subtracted from it. Thus, no signal is present.

## Slide 2
COMPENSATION
FITC
FITC -
PE
PE -
FITC
-
-
+
+
gain=1
gain=1
detector
detector
PE
PE
FL1
FL2
FL1-%FL2
Notes: Similarily, if particles stained with PE are used a proportion of the PE detector output can be directed to the FITC op amp (top) exactly equal to the signal from the FITC detector, resulting in no output (FL1-%FL2). Any number of these compensation circuits can be appropriately combined to provide spectral compensation for any number of colors.

## Slide 3
Uncompensated vs Compensated
FL1
FL2
Notes: In A, the amount of fluorescence emission detected by the PE detector and the amount of PE fluorescence detected by the FITC detector is shown. After compensation, B, the particles show only their appropriate fluorescence emission.

## Slide 4
FL1
FL2
COMPENSATION IS INTENSITY DEPENDENT
uncompensated
partially compensated
fully compensated
Notes: It is an absolute requirement that particles used for compensation be at least as bright as the antibody staining. If this condition is not met, antibodies whose fluorescence is brighter than the particles used for compensation will not be properly compensated.

## Slide 5
COMPENSATE INSTRUMENT USING STAINED CELLS
1. Adjust PMT voltages using unstained cells
2. Adjust compensation for each fluorochrome
Notes: We recommend instrument compensation using stained cells because appropriate particles are not currently available for the fluorochromes that are available.

## Slide 6
0
256
512
768
1024
030695001
FSC ->
0
256
512
768
1024
030695001
FSC ->
Lymphogated Cells for Four Color
Single Laser Compensation
A
B
FSC
FSC
SSC
SSC
Notes: Lymphocytes are like microspheres in that antibodies can be chosen that will provide uniform fluorescence. Furthermore, microspheres are not available with every fluorochrome. Using a sample of unstained cells, set FSC and SSC to display cells properly (A). Establish a live gate around lymphocytes (B). Make sure all compensations are set at 0%.

## Slide 7
Notes: Set PMT voltages for all fluorochromes so that cells are contained inside the 101 x 101 quadrant of each bivariate plot but make sure FL3 is 15 volts higher than FL4.

## Slide 8
Notes: Change cells to those stained with FITC and adjust FL2 - %FL1 and FL4 - %FL1 so that cells are within their appropriate quadrants as shown.

## Slide 9
Notes: The order of compensation adjustment is critical. Do not deviate or overcompensation of some fluorochromes will occur. Put the PE stained cells on and adjust FL1 - %FL2, FL3 - %FL2 and FL4 - %FL2 so that cells are inside their appropriate quadrants as shown.

## Slide 10
Notes: Change cells to those stained with the PE-TxRd tandem complex and adjust FL2 - %FL4 and FL3 - %FL4 as shown.

## Slide 11
Notes: Change cells to those stained with PE-CY5 tandem complex or PerCP and adjust FL2 - %FL3 and FL4 - %FL3 as shown.

## Slide 12
FSC
SSC
R1
R2
R3
R4
FITC-CD45
PE-CD4
PE-CY5-CD8
PE-CD4
R4
R5
R6
PE-CY5-CD8
APC-CD8
R8
R9
R10
COMPENSATION UP TO FOUR COLORS WITH TWO LASERS
Notes: Any number of parameter combinations can be properly compensated using cells stained with a single antibody and then mixing them together for adjustment. In this example, human blood was used. Regions are created to position proper compensation of cells. Cells stained with each color must fall within its appropriate region. Each day a new compensation standard is prepared. By acquiring both the newly prepared standard and the previous day’s standard each day (gated on R1), it can be verified that both the instrument and the new standard are functioning properly because cells from both are in their correct regions. If the new standard contains one or more cell populations that fall outside their region, but the old one does not, the standard has been improperly made. If, however, both standards fall outside their respective regions then the instrument requires service.

## Slide 13
Notes: When a single laser is used for excitation, FITC, PE, PE-CY5 or PerCP conjugated antibodies can be used with a PE-TR conjugated reagent. There are, however, two problems with this approach. First, the PE-TR emission significantly overlaps both PE (because of PE leakage from the complex) and TR, making compensation problematic. Secondly, there is a very limited selection of PE-TR conjugated antibodies available, but using a biotin conjugated antibody with RED613 avidin provides increased flexability. To verify proper compensation, cells must be stained with a unique antibody combination that will produce the following distinct populations: 1) FITC+PE-, 2) PE+TC+FC-, 3) PE-TC+FC+, 4) PE-TC-FC+ (TC= third color, PE-CY5, or PerCP and FC= fourth color, ECD or RED613). A suitable antibody combination to provide these populations is FITC-CD3, PE-CD8,TC-CD2 and FC-CD4. Cells are stained with this combination in a single tube. This figure shows the uncompensated data for cells stained with this combination. Because of the increased dimensionality of the data, compensation is actually unnecessary as each population is distinctly resolved..

## Slide 14
R1
R2
R3
R4
R5
R6
Notes: Software compensation offers a new and powerful means of retrospectively compensating data. Cells stained with each antibody separately are acquired to produce files that are used to develop the compensation matrix. This only needs to be done once. The matrix is then applied to the verification antibody combination run daily. The properly compensated cells produce the pattern shown here. Regions and quadrant markers are established so that each cell population is found in its proper region.

## Slide 15
Notes: While software compensation has the advantage of allowing for precise retrospective adjustment of compensation and is very easy, electronic instrument compensation, when properly performed, can produce the same results. Can you tell the difference? (clue: compare cells along the axis.)

## Slide 16
R3
R3
A
C
B
D
Notes: A improperly compensated instrument almost always involves the PE, PE-TR, PE-CY5 antibodies. Over or under compensation will always result in artifactual data, so verification of proper compensation must be performed for each experiment, without exception. In A, FL2-%FL3 is undercompensated so that not all the CD8+ cells express CD2, but this is an artifactual result of improper compensation, as shown in B. In C, undercompensation of FL4-%FL2 causes some CD8+ cells to appear as though they express CD4, but this is also an artifactual result, as shown in D. Four color single laser compensation is more difficult than four color two laser compensation because of emission spectra overlap.